## Modbus Register
Category | Name | Address | Size | Type | Variable | Units | Expected Data | Purpose/Function | Notes
Controller | Boot | 1 | - | Discrete |  | - |  | Indicates when Sentro X is in Boot Mode
Controller | Sensing | 2 | - | Discrete |  | - |  | Indicates when Sentro X is in Sensing Mode
Controller | Maintenance | 3 | - | Discrete |  | - |  | Indicates when Sentro X is in Maintenance Mode
Controller | Warning | 4 | - | Discrete |  | - |  | Indicates when Sentro X is in Warning Mode
Controller | Alarm | 5 | - | Discrete |  | - |  | Indicates when Sentro X is in Alarm Mode
Controller | Fault | 6 | - | Discrete |  | - |  | Indicates when Sentro X is in Fault Mode
Controller | Overrange | 7 | - | Discrete |  | - |  | Indicates when Sentro X is in Overrange Mode
Controller | Underrange | 8 | - | Discrete |  | - |  | Indicates when Sentro X is in Underrange
Controller | Pellistor Lockout | 9 | - | Discrete |  | - |  | Indicates when Sentro X is in Pellistor Lockout
Controller | Firmware Updating | 10 | - | Discrete |  | - |  | Sentro X firmware is being updated
Controller | Firmware Updating Complete | 11 | - | Discrete |  | - |  | Sentro X firmware has been successfully updated
Controller | Delta Bloc Present | 12 | - | Discrete |  | - |  | The Delta Bloc is connected in place of a gBloc
Controller | Passcode | 1 | 1 | Holding | unsigned | - | 0000 - 9999 | Set a passcode for the device via Modbus
Controller | Passode Enable | 1 | - | Coil | - | - | 0 = Disabled, 1 = Enabled | Enable or disable the passcode
Controller | Serial number | 1 | 15 | Input | ASCII[15] | - | AA01234567 |  | Each register contains two characters
Controller | TX Number | 16 | 15 | Input | ASCII[15] | - | TX6310.00.00.00.00 |  | Breakdown here,
Controller | Firmware Version | 31 | 4 | Input | unsigned[4] | - | 1.0.0.0
Controller | Variant | 35 | 1 | Input | unsigned | - | 0 | Currently unused
gBloc | Serial number | 36 | 15 | Input | ASCII[15] |  | AA01234567 | gBloc Serial Number | Each register contains two characters
gBloc | TX Number | 51 | 15 | Input | ASCII[15] |  | TX6310.00.000 | gBloc TX Number | TX6304 for the Delta Bloc, Each register contains two characters
gBloc | Firmware Version | 66 | 4 | Input | unsigned[4] |  | 1.0.0.0 | Firmware version of the gBloc
gBloc | Variant | 70 | 1 | Input | unsigned |  | - | Currently unused
gBloc | Gas String Short | 71 | 4 | Input | ASCII[4] | - | CO, CO2, CH4 etc. | gBloc short gas string | Definitions here
gBloc | Gas String Long | 75 | 15 | Input | ASCII[15] | - | Carbon Monoxide, Carbon Dioxide, Methane etc. | gBloc long gas string | Definitions here
gBloc | Bloc Cell Type Short | 90 | 4 | Input | ASCII[4] | - | eChem, Pel, IR etc | gBloc cell type short string | Definitions here
gBloc | Bloc Cell Type Long | 94 | 15 | Input | ASCII[15] | - | Electrochemical, Pellistor, Infrared etc | gBloc cell type long string | Definitions here
gBloc | Range (Lower) | 109 | 2 | Input | float | - | 1.00, 10.0, 100, 1000 | Lower range of the gBloc | The number format here will depend on the format specified by the gBloc
gBloc | Range (Upper) | 111 | 2 | Input | float | - | 1.00, 10.0, 100, 1000 | Upper range of the gBloc | The number format here will depend on the format specified by the gBloc
gBloc | Gas Reading Units | 113 | 1 | Input | unsigned | - | 0 - 7 | Gas concentration display units | See here
gBloc | Gas Reading Decimal Places | 114 | 1 | Input | unsigned | - | 0 - 3 | Gas concentration number of decimal places | Precision of the gas reading
gBloc | Gas Reading Whole Numbers | 115 | 1 | Input | unsigned | - | 0 - 3 | Gas concentration number of whole numbers | Maximum number of whole numbers to be displayed
gBloc | Warmup time | 116 | 1 | Input | unsigned | seconds | 0-X | Warmup time of the gBloc
gBloc | Warmup timer live | 117 | 1 | Input | unsigned | seconds | 0-X | Current warmup time
gBloc | LEL Scale | 13 | - | Discrete |  | - | 0 = Non LEL gBloc variant | Indicates if the gBloc scale is an LEL variant
gBloc | LEL Value (%) | 118 | 2 | Input | float |  | 0.0-100.0 | LEL value in %v/v
Setpoint 1 | Setpoint 1 Enable / Disable | 2 | - | Coil | - | - | 0 = Disabled, 1 = Enabled | Enable or disable the setpoint
Setpoint 1 | Setpoint 1 Mode | 3 | - | Coil | - | - | 0 = Under, 1 = Over | Change the trigger direction of the setpoint
Setpoint 1 | Setpoint 1 Latch | 4 | - | Coil | - | - | 0 = No latch, 1 = Latch | Change the setpoint to latching or non-latching
Setpoint 1 | Setpoint 1 Assign Alarm | 5 | - | Coil | - | - | 0 = Warning, 1 = Alarm | Assign an alarm severity
Setpoint 1 | Setpoint 1 Level | 2 | 2 | Holding | float | variable | 0 - Range of gBloc |  | The number format here will depend on the format specified by the gBloc
Setpoint 1 | Setpoint 1 Hysteresis | 4 | 1 | Holding | unsigned | % | 0 - 99 |  | 10% hysteresis on a 1%v/v setpoint is 0.9%v/v
Setpoint 1 | Setpoint 1 Assign Relay | 5 | 1 | Holding | unsigned | - | 0000 - 1111 (RA1/RA2/RA3/RA4) | Which relays are assigned to the setpoint
Setpoint 1 | Setpoint 1 Active | 14 | - | Discrete |  | - | 0 = Not Active, 1 = Active
Setpoint 1 | Setpoint 1 Reset | 23 | - | Coil | - | - | 0 = Default, 1 = Reset requested | Send 1 to reset the active alarm (if resettable)
Setpoint 2 | Setpoint 2 Enable / Disable | 6 | - | Coil | - | - | 0 = Disabled, 1 = Enabled | Enable or disable the setpoint
Setpoint 2 | Setpoint 2 Mode | 7 | - | Coil | - | - | 0 = Under, 1 = Over | Change the trigger direction of the setpoint
Setpoint 2 | Setpoint 2 Latch | 8 | - | Coil | - | - | 0 = No latch, 1 = Latch | Change the setpoint to latching or non-latching
Setpoint 2 | Setpoint 2 Assign Alarm | 9 | - | Coil | - | - | 0 = Warning, 1 = Alarm | Assign an alarm severity
Setpoint 2 | Setpoint 2 Level | 6 | 2 | Holding | float | variable | 0 - Range of gBloc |  | The number format here will depend on the format specified by the gBloc
Setpoint 2 | Setpoint 2 Hysteresis | 8 | 1 | Holding | unsigned | % | 0 - 99 |  | 10% hysteresis on a 1%v/v setpoint is 0.9%v/v
Setpoint 2 | Setpoint 2 Assign Relay | 9 | 1 | Holding | unsigned | - | 0000 - 1111 (RA1/RA2/RA3/RA4) | Which relays are assigned to the setpoint
Setpoint 2 | Setpoint 2 Active | 15 | - | Discrete |  | - | 0 = Not Active, 1 = Active
Setpoint 2 | Setpoint 2 Reset | 24 | - | Coil | - | - | 0 = Default, 1 = Reset requested | Send 1 to reset the active alarm (if resettable)
Setpoint 3 | Setpoint 3 Enable / Disable | 10 | - | Coil | - | - | 0 = Disabled, 1 = Enabled | Enable or disable the setpoint
Setpoint 3 | Setpoint 3 Mode | 11 | - | Coil | - | - | 0 = Under, 1 = Over | Change the trigger direction of the setpoint
Setpoint 3 | Setpoint 3 Latch | 12 | - | Coil | - | - | 0 = No latch, 1 = Latch | Change the setpoint to latching or non-latching
Setpoint 3 | Setpoint 3 Assign Alarm | 13 | - | Coil | - | - | 0 = Warning, 1 = Alarm | Assign an alarm severity
Setpoint 3 | Setpoint 3 Level | 10 | 2 | Holding | float | variable | 0 - Range of gBloc |  | The number format here will depend on the format specified by the gBloc
Setpoint 3 | Setpoint 3 Hysteresis | 12 | 1 | Holding | unsigned | % | 0 - 99 |  | 10% hysteresis on a 1%v/v setpoint is 0.9%v/v
Setpoint 3 | Setpoint 3 Assign Relay | 13 | 1 | Holding | unsigned | - | 0000 - 1111 (RA1/RA2/RA3/RA4) | Which relays are assigned to the setpoint
Setpoint 3 | Setpoint 3 Active | 16 | - | Discrete |  | - | 0 = Not Active, 1 = Active
Setpoint 3 | Setpoint 3 Reset | 24 | - | Coil | - | - | 0 = Default, 1 = Reset requested | Send 1 to reset the active alarm (if resettable)
TWA | TWA Enable | 14 | - | Coil | - | - | 0 = Disabled, 1 = Enabled | Enable or disable the TWA setpoint | Enabling/Disabling clears the TWA value and period registers
TWA | TWA Limit | 14 | 2 | Holding | float | variable |  | Set the TWA alarm value | The number format here will depend on the format specified by the gBloc
TWA | TWA Time (Hours) | 16 | 1 | Holding | unsigned | hours | 8-12 hours | Change the TWA period | Changing the period NOT recalculate the level for you. It must be done manually
TWA | TWA Assign Alarm | 15 | - | Coil | - | - | 0 = Warning, 1 = Alarm | Assign an alarm severity
TWA | TWA Assign Relay | 17 | 1 | Holding |  |  | 0000 - 1111 (RA1/RA2/RA3/RA4) | Which relays are assigned to the TWA setpoint
TWA | TWA Reset | 16 | - | Coil | - | - | 0 = Default, 1 = Set to zero | Reset the TWA value and period to zero
TWA | TWA Active | 17 | - | Discrete |  | - | 0 = Not Active, 1 = Active
TWA | TWA Period Elapsed (Hours) | 120 | 1 | Input | unsigned | hours | 0-12
TWA | TWA Period Elapsed (Minutes) | 121 | 1 | Input | unsigned | minutes | 0-59
TWA | TWA Period Elapsed (Seconds) | 122 | 1 | Input | unsigned | seconds | 0-59
STEL | STEL Enable / Disable | 17 | - | Coil | - | - | 0 = Disabled, 1 = Enabled | Enable or disable the STEL setpoint | Enabling/Disabling clears the STEL value register
STEL | STEL Limit | 18 | 2 | Holding | float | variable | 0 - Range of gBloc | Set the STEL alarm value | The number format here will depend on the format specified by the gBloc
STEL | STEL Latch | 18 | - | Coil | - | - | 0 = Non-Latching, 1 = Latching | Change the setpoint to latching or non-latching
STEL | STEL Assign Alarm | 19 | - | Coil | - | - | 0 = Warning, 1 = Alarm | Assign an alarm severity
STEL | STEL Assign Relay | 20 | 1 | Holding | unsigned int |  | 0000 - 1111 (RA1/RA2/RA3/RA4) | Which relays are assigned to the STEL setpoint
STEL | STEL Reset | 20 | - | Coil | - | - | 0 = Default, 1 = Set to zero | Reset the STEL value and persiod to zero
STEL | STEL Active | 18 | - | Discrete |  | - | 0 = Not Active, 1 = Active
Faults | Faults Assign Relay | 21 | 1 | Holding | unsigned | - | 0000 - 1111 (RA1/RA2/RA3/RA4) | Which relays are assigned to the Faults state
Faults | Fault Code | 123 | 1 | Input | unsigned | - | 0-8 | Provide fault information | Definitions here (more may be added later)
Overrange | Overrange Assign Relay | 22 | 1 | Holding | unsigned | - | 0000 - 1111 (RA1/RA2/RA3/RA4) | Which relays are assigned to the overrange setpoint | This register is shared with Pelistor protect overrange also
Readings | Gas | 124 | 2 | Input | float | - | - | Gas concentration value (Variable units) | -10,000 default value
Readings | Temperature | 126 | 2 | Input | float | °C | 0 |  | -10,000 default value
Readings | Pressure | 128 | 2 | Input | float | mBar | 0 |  | -10,000 default value
Readings | Humidity | 130 | 2 | Input | float | % | 0 |  | -10,000 default value
Readings | Supply voltage | 132 | 2 | Input | float | V | Vsupply
Readings | TWA | 134 | 2 | Input | float | - |  | TWA value (variable units) | The number format here will depend on the format specified by the gBloc
Readings | STEL | 136 | 2 | Input | float | - |  | STEL value (variable units) | The number format here will depend on the format specified by the gBloc
Maintenance | Zero Test completed | 19 | - | Discrete |  | - |  | Test has finished
Maintenance | Zero Test completed with adjustment | 20 | - | Discrete |  | - |  | Test has finished, adjustment was made
Maintenance | Zero Test adjusted by | 138 | 2 | Input | float | variable |  | ppm adjusted to zero the sensor | Do we want % versions of this registers also?
Maintenance | Span Test completed | 21 | - | Discrete |  | - |  | Test has finished
Maintenance | Span Test completed with adjustment | 22 | - | Discrete |  | - |  | Test has finished, adjustment was made
Maintenance | Span Test adjusted by | 140 | 2 | Input | float | variable |  | ppm adjusted to get the sensor to the target gas
Maintenance | Span Test target gas | 142 | 2 | Input | float | variable |  | Last target gas used on a span test completed with adjust
Maintenance | Response Time Test completed | 23 | - | Discrete |  | -
Maintenance | Response Time | 144 | 2 | Input | float | seconds | e.g. 21.6s | Last T90 response time test time in seconds
Maintenance | Linearity completed | 24 | - | Discrete |  | - |  | Linearity test completed flag
Maintenance | Linearity Check 1 target gas | 146 | 2 | Input | float | variable |  | Target gas used for the first point of linearity check | The number format here will depend on the format specified by the gBloc
Maintenance | Linearity Check 1 result | 148 | 2 | Input | float | variable
Maintenance | Linearity Check 2 target gas | 150 | 2 | Input | float | variable
Maintenance | Linearity Check 2 result | 152 | 2 | Input | float | variable
Maintenance | Linearity Check 3 target gas | 154 | 2 | Input | float | variable
Maintenance | Linearity Check 3 result | 156 | 2 | Input | float | variable
Maintenance | Linearity Check 4 target gas | 158 | 2 | Input | float | variable
Maintenance | Linearity Check 4 result | 160 | 2 | Input | float | variable
Maintenance | Linearity Check 5 target gas | 162 | 2 | Input | float | variable
Maintenance | Linearity Check 5 result | 164 | 2 | Input | float | variable
Maintenance | Linearity Check 6 target gas | 166 | 2 | Input | float | variable
Maintenance | Linearity Check 6 result | 168 | 2 | Input | float | variable
Maintenance | Telemetry Test completed | 25 | - | Discrete | - | - |  | Telemetry test completed flag
Maintenance | Telemetry Test Flag | 26 | - | Discrete | - | -
Maintenance | Dynamic Range Test completed | 27 | - | Discrete | - | - |  | Dynamic range test completed flag
Maintenance | Indicators Test completed | 28 | - | Discrete | - | - |  | Audio and visual alarm test completed flag
Hardware | I/O Card Slot 1 | 170 | 1 | Input | unsigned |  | 0-1 | 0 = Not Fitted, 1 = Relay Card
Hardware | I/O Card Slot 2 | 171 | 1 | Input | unsigned |  | 0-2 | 0 = Not Fitted, 1 = Relay Card, 2 = Ethernet Card
Hardware | AV Alarm Fitted | 29 | - | Discrete |  | - |  | 0 = Not Fitted, 1 = Fitted
Hardware | Remote Head Fitted | 30 | - | Discrete |  | - |  | 0 = Not Fitted, 1 = Fitted
Hardware | Beacon Assign | 23 | 1 | Holding | unsigned | - | 0000 - 1111 (Warning/Alarm/Fault/Overrange)
Hardware | Beacon Status | 31 | - | Discrete |  | - | 0 = Off, 1 = On
Hardware | Sounder Assign | 24 | 1 | Holding | unsigned | - | 0000 - 1111 (Warning/Alarm/Fault/Overrange)
Hardware | Sounder Status | 32 | - | Discrete |  | - | 0 = Off, 1 = On
Hardware | Sounder Mute | 21 | - | Coil | - | - | 0 = Sounder not muted, 1 = Sounder muted | Mute the sounder if it is active | Register is reset when the sounder on condition is reset (Fault/Alarm is cleared etc)
Hardware | Relay A1 | 33 | - | Discrete |  | - | 0 = Inactive, 1 = Active | Status of the relay | Active means de-energised
Hardware | Relay A2 | 34 | - | Discrete |  | - | 0 = Inactive, 1 = Active | Status of the relay
Hardware | Relay B1 | 35 | - | Discrete |  | - | 0 = Inactive, 1 = Active | Status of the relay
Hardware | Relay B2 | 36 | - | Discrete |  | - | 0 = Inactive, 1 = Active | Status of the relay
Ethernet | MAC Address | 172 | 6 | Input | unsigned[6] |  |  | MAC address of the relay card
Ethernet | IP Address | 178 | 4 | Input | unsigned[4]
Ethernet | Subnet Mask | 182 | 4 | Input | unsigned[4]
Ethernet | Gateway | 186 | 4 | Input | unsigned[4]
Ethernet | Config | 37 | - | Discrete |  | - | 0 = Static, 1 = Dynamic | IP, Subnet Mask & Gateway are automattically fetched on Dynamic
Analogue Out | Type | 38 | - | Discrete |  | - | 0 = Current, 1 = Voltage | Analogue output type configuration
Debug | ECC Single Error Count | 190 | 4 | Input | unsinged | - |  | Counts how many single error ECC corrections have taken place
Debug | Scada String Compatibility Flag | 22 | - | Coil | - | - | 0 = Disabled, 1 = Enabled | Changes the way strings are copied to the input register
Debug | Uptime Counter | 194 | 4 | Input | unsigned | - | Seconds value since boot
Trolex Only | Manufacturing Register Enable | 420000 | 1 | Holding | unsigned | - | Passcode |  | Must write to the save register to save the written serial number and tx number. These registers are only accessable using the WRITE_SINGLE_HOLDING_REG (Function 06). So must be written to one character at a time. MAC address writing & saving is separate to the system info save.
Trolex Only | Sentro X Serial Number | 420001 | 30 | Holding | char | - | Ascii characters, 1 character per address | Sentro X configurator to write serial number
Trolex Only | Sentro X TX number | 420031 | 30 | Holding | char | - | Ascii characters, 1 character per address | Sentro X configurator to write tx number
Trolex Only | Save System Info | 420061 | 1 | Holding | unsigned | - | Write 1 to save
Trolex Only | Sentro X Ethernet Card MAC address | 420062 | 6 | Holding | unsigned | - | MAC address | Backup for writing the MAC address to the ethernet card
Trolex Only | Save MAC address | 420068 | 1 | Holding | unsigned | - | Write 1 to save | Save the MAC address to the Ethernet card
F01 Coil = Read/Write flag
F02 Discrete = Read-Only flag
F03 Holding = Read/Write value
F04 Input = Read-Only value

## Gas Strings
Short | Long
NH3 | Ammonia
CO2 | Carbon Dioxide
CO | Carbon Monoxide
H2 | Hydrogen
H2S | Hydrogen Sulphide
CH4 | Methane
NO | Nitric Oxide
NO2 | Nitrogen Dioxide
O2 | Oxygen
SO2 | Sulphur Dioxide

## Gas Units
Gas Scale | Value
None | 0
PPM | 1
%v/v | 2
PPB | 3
%LEL | 4
Milliamp | 5
Volts | 6
% | 7

## Sensor Type
Sensor Type | Short String | Long String
None | NONE | NONE
Electrochemical | eChem | Electrochemical
Infrared | IR | Infrared
Pellistor | Pel | Pellistor
MEMs | MEMs | MicroElectroMechanical
Potentiometer | Delta | Delta
Catilitic | Cat | Catalitic
MOS | MOS | Metal Oxide Semiconductor

## Fault Codes
Fault Code | On Screen Text | Description
0 | None | No fault code
1 | Underrange Lockout | The gas reading has dropped below -10% the range of the cell
2 | UI Timeout | The UI has been abandoned on a certain screen
3 | gBloc Timeout | Failed to insert the gBloc in time
4 | gBloc Connection | Connection to the gBloc has been lost
5 | Undervoltage | Supply voltage rail has dropped below 9.0v
6 | Boot IO Module | Expansion modules are in the incorrect position
7 | Boot System Info | System Info has been incorrectly read from the flash (flash error)
8 | Boot System Settings | System settings has been incorrectly read from the flash (flash error)

## TX Number
TX6310 | .AA | Cert | .BB | Cover | .CC | AV Alarm | .DD | Extention Cards
| .00 | General Purpose | .01 | Internal gas sensor | .00 | Not fitted | .00 | Not fitted
| .01 | ATEX Grp 1 | .02 | Remote gas sensor (cover + rBloc) | .01 | Fitted | .01 | 1 x Dual relay card
| .06 | IECEx Australia Grp 1 |  |  |  |  | .02 | 2 x Dual relay card
[7 empty] | .03 | 1 x Ethernet card
[7 empty] | .04 | 1 x Dual relay + Ethernet